Quality Control Report – November 2022

The quality control report is generated, using Applied Biosystems Sequence Scanner v1.0, from the results of the sequencing of control samples that are produced within the Core DNA Sequencing Unit. The controls are pGEM® DNA (Applied Biosystems) that are sequenced with the M13F primer (IDT).

The resultant quality control report is saved as a pdf file and is sent together with this letter that gives explanations as to the possible reasons why some of the reactions have failed. The pdf file of the quality control report gives a summery of the quality of the trace score as well as that of the contiguous read length (CRL) of the control samples. This form is produced at the end of every month.

Failed Samples:

| Trace File Name | Trace Score | Comment |
| --- | --- | --- |
| A09_Control_442_09_11_2022_1.ab1 | 34 | Spike |
| G11_Control_444_10_11_2022_1.ab1 | 31 | Sephadex or injection error |
| F12_Control_468_28_11_2022_1.ab1 | NA | Sephadex error |
| H09_Control_467_28_11_2022_1.ab1 | 19 | Sephadex error |
| G12_Control_476_30_11_2022_1.ab1 | 13 | Sephadex error |
| G12_Control_478_30_11_2022_1.ab1 | 10 | Sephadex error |
| H12_Control_477_30_11_2022_1.ab1 | 10 | Sephadex error |
| Sephadex error experienced on Tecan when alternative 50 ul tips were used that was longer and pierced the sephadex during transfer sporadically. Changed back to the original tips. |  |  |