Quality Control Report – April 2024

The quality control report is generated, using Applied Biosystems Sequence Scanner v1.0, from the results of the sequencing of control samples that are produced within the Core DNA Sequencing Unit. The controls are pGEM® DNA (Applied Biosystems) that are sequenced with the M13F primer (IDT).

The resultant quality control report is saved as a pdf file and is sent together with this letter that gives explanations as to the possible reasons why some of the reactions have failed. The pdf file of the quality control report gives a summery of the quality of the trace score as well as that of the contiguous read length (CRL) of the control samples. This form is produced at the end of every month.

Failed Samples:

| Trace File Name | Trace Score | Comment |
| --- | --- | --- |
| H08_Control_96_16_04_2024_1.ab1 | 20 | Sephadex or electrophoresis error as a short sequence produced which tapers of quickly. |
| H03_Control_107_25_04_2024_1.ab1 | 36 | Sephadex or electrophoresis error as a short sequence produced which tapers of quickly. |
| G12_Control_110_26_04_2024_1.ab1 | 34 | Sephadex or electrophoresis error as a short sequence produced which tapers of quickly. |